## Sheet1
Method_Name | Method_ID | Method_Type | Method_Description | Method_Instrument | Method_Lab
Frozen_Flag_Code | FLAG_FRZ_001 | flag | Vials were frozen (-20 degC) prior to analysis. Freezing may affect SpC data. ICP data. and/or IC data. | NA | Pacific Northwest National Laboratory
Frozen_Flag_Code | FLAG_FRZ_002 | flag | Vials were frozen during shipment. Freezing may affect IC data and or SpC Data.
Ammonia_Method | AMM_001 | MethodID_Analysis | Ammonium concentrations in water samples were determined photometrically. Absorbances were measured using a Shimadzu UV mini 1240 at 630 nanometers. A standard suite range of 5-200 µM was used. | Shimadzu UV mini 1240 | Pacific Northwest National Laboratory
Anion_Software_Method | ANIONSFTW_000 | MethodID_Analysis | Anions not measured | NA | Pacific Northwest National Laboratory
Anion_Software_Method | ANIONSFTW_001 | MethodID_Analysis | Ions: F. Cl. NO2. SO4. NO3; Std concentration ranges (mg/L): (F: 0.2-10). (Cl: 0.16-80). (NO2: 0.04-20). (SO4: 0.16-80). (NO3: 0.12-60); Dionex ICS-2000. Ion Pack AS18. prod#: 060549. serial#: 180602097. elution buffer: 22 mM KOH; suppressor current: 57 mA | Dionex ICS-2000 | Pacific Northwest National Laboratory
Anion_Software_Method | ANIONSFTW_002 | MethodID_Analysis | Ions: F. Cl. NO2. SO4. NO3; Std concentration ranges (mg/L): (F: 0.2-10). (Cl: 0.16-80). (NO2: 0.04-20). (SO4: 0.16-80). (NO3: 0.12-60). (Br: 0.03-15); Dionex ICS-2000. Ion Pack AS18. prod#: 060549. serial#: 180602097. elution buffer: 22 mM KOH; suppressor current: 57 mA | Dionex ICS-2000 | Pacific Northwest National Laboratory
Anion_Software_Method | ANIONSFTW_003 | MethodID_Analysis | Need to fill in for Sequim
DIC_Software_Method | DICRUN_001 | MethodID_Analysis | Inorganic carbon (IC) was determined by using a Shimadzu combustion carbon analyzer TOC-L CSH/CSN E100V with a manual injection block. 50 uL samples were injected into 25 % phosphoric acid (ACS grade Sigma Aldrich#438081-500mL) at ambient temperature. Inorganic carbon standard was made from sodium bicarbonate solid (Nacalia Tesque lot M7A5839) and sodium carbonate solid (Nacalia Tesque lot M7A6147). The calibration range for inorganic carbon was 2.5 to 50 ppm C. | Shimadzu TOC-L | Pacific Northwest National Laboratory
FTICR_Analysis_Method | ICRRUN_001 | MethodID_Analysis | Ion accumulation time (IAT) = 0.05 | 12 Tesla (12T) Bruker SolariX FTICR-MS | Environmental Molecular Sciences Laboratory/Pacific Northwest National Laboratory
FTICR_Analysis_Method | ICRRUN_002 | MethodID_Analysis | Ion accumulation time (IAT) = 0.1 | 12 Tesla (12T) Bruker SolariX FTICR-MS | Environmental Molecular Sciences Laboratory/Pacific Northwest National Laboratory
FTICR_Analysis_Method | ICRRUN_003 | MethodID_Analysis | Ion accumulation time (IAT) = 0.08 and 0.1 | 12 Tesla (12T) Bruker SolariX FTICR-MS | Environmental Molecular Sciences Laboratory/Pacific Northwest National Laboratory
FTICR_Analysis_Method | ICRRUN_004 | MethodID_Analysis | Ion accumulation time (IAT) = 0.08 | 12 Tesla (12T) Bruker SolariX FTICR-MS | Environmental Molecular Sciences Laboratory/Pacific Northwest National Laboratory
FTICR_Analysis_Method | ICRRUN_005 | MethodID_Analysis | Ion accumulation time (IAT) = 0.3 | 12 Tesla (12T) Bruker SolariX FTICR-MS | Environmental Molecular Sciences Laboratory/Pacific Northwest National Laboratory
FTICR_Analysis_Method | ICRRUN_006 | MethodID_Analysis | Ion accumulation time (IAT) = 0.08 and 0.05 | 12 Tesla (12T) Bruker SolariX FTICR-MS | Environmental Molecular Sciences Laboratory/Pacific Northwest National Laboratory
FTICR_Analysis_Method | ICRRUN_007 | MethodID_Analysis | Ion accumulation time (IAT) = 0.08. The sample did not calibrate well during testing (although only uncalibrated data are provided in data package) and did not meet QA/QC standards. | 12 Tesla (12T) Bruker SolariX FTICR-MS | Environmental Molecular Sciences Laboratory/Pacific Northwest National Laboratory
FTICR_Batch_Method | ICRBAT_001 | MethodID_Analysis | The samples were run together the week of 4 November 2019 | NA | Pacific Northwest National Laboratory
FTICR_Batch_Method | ICRBAT_002 | MethodID_Analysis | The samples were run together the week of 2 December 2019 | NA | Pacific Northwest National Laboratory
FTICR_Batch_Method | ICRBAT_003 | MethodID_Analysis | The samples were run together the week of 9 December 2019 | NA | Pacific Northwest National Laboratory
FTICR_Batch_Method | ICRBAT_004 | MethodID_Analysis | The samples were run together the week of 1 August 2019 | NA | Pacific Northwest National Laboratory
FTICR_Batch_Method | ICRBAT_005 | MethodID_Analysis | The samples were run together the week of 2 March 2020 | NA | Pacific Northwest National Laboratory
FTICR_Batch_Method | ICRBAT_006 | MethodID_Analysis | The samples were run togeher the week of 21 September 2020 | NA | Pacific Northwest National Laboratory
FTICR_Batch_Method | ICRBAT_007 | MethodID_Analysis | The samples were run together the week of 8 July 2020. | NA | Pacific Northwest National Laboratory
FTICR_Batch_Method | ICRBAT_008 | MethodID_Analysis | The samples were re-run together the week of 27 August 2020. They had already been previously run for ICR. | NA | Pacific Northwest National Laboratory
FTICR_Batch_Method | ICRBAT_009 | MethodID_Analysis | The samples were run together the week of 17 March 2020 | NA | Pacific Northwest National Laboratory
FTICR_Format_Method | ICRFORM_000 | MethodID_Analysis | The FTICR-MS data did not pass QAQC. EMSL did not release the data for use. | NA | Pacific Northwest National Laboratory
FTICR_Format_Method | ICRFORM_001 | MethodID_Analysis | Columns in FTICR Output are as follows: Mass (measured ionic mass; m/z). Columns C. H. O. N. C13. S. P. Na (atoms in the analyzed molecule. Idenitified by an algorithm in Kujawinski and Behn 2006 and modified as described in Tfaily et al. 2017.) El_Comp (function group(s) that are known. NeutralMass (zero-charge or neutral molecular mass. Calculated from observed ionic mass by balancing charge with protons or electrons). Error_ppm (parts per million error calculated as [expected molecular mass - measured molecular mass]/expected molecular mass * 1e6). Candidates (possible known molecules the analyzed molecule could be. Assigned by checking against a known database. Value of -1 when molecule was not checked against database). All remaining columns identified by Kit ID and Rep number (spectra for each sample where the numerical value is the peak intensity at the particular molecular mass). | NA | Pacific Northwest National Laboratory
ICP_Software_Method | ICPSFTW_000 | MethodID_Analysis | ICP not measured | NA | Pacific Northwest National Laboratory
ICP_Software_Method | ICPSFTW_001 | MethodID_Analysis | Elements: P. In. Fe. Mg. Ca. K. Na; Std concentration ranges (ug/L): 5 to 4000 for all except Fe. 0.5-400 for Fe; Perkin Elmer Optima 2100 DV | Perkin Elmer Optima 2100 DV | Pacific Northwest National Laboratory
InorgC_Software_Method | INORGCSFTW_000 | MethodID_Analysis | Inorganic carbon not measured | NA | Pacific Northwest National Laboratory
InorgC_Software_Method | INORGCSFTW_001 | MethodID_Analysis | Best 4 of 5 injections; injection vol: 50 uL; furnace temp: 680 degC; std concentration ranges: 10-30 mg/L ; Shimadzu TOC-V | Shimadzu TOC-V | Pacific Northwest National Laboratory
Isotope_Software_Method | ISOTO_000 | MethodID_Analysis | Stable isotopes not measured | NA | Pacific Northwest National Laboratory
Isotope_Software_Method | ISOTO_001 | MethodID_Analysis | Water samples were run on a Los Gatos Research water analyzer for stable isotopes (del-18-O and del-2-H). 8 replicate injections of each sample - with the first 4 discarded to ensure no carryover. 5 Los Gatos Research water standards were run before and after each sample set for isotopic calibration via a point slope correction. Also ran an aliquot of previously calibrated (by USGS table isotope laboratory in Reston VA) water as a check on system performance. | Los Gatos Research water analyzer | Environmental Molecular Sciences Laboratory/Pacific Northwest National Laboratory
NPOC_RERUN | NPOCRERUN_001 | MethodID_Analysis | Sample was re-run because it fell outside of the calibration curve. An additional dilution was made from the leftover sample from the first try. |  | Pacific Northwest National Laboratory
NPOC_Software_Method | NPOC_000 | MethodID_Analysis | NPOC not measured | NA | Pacific Northwest National Laboratory
NPOC_Software_Method | NPOC_001 | MethodID_Analysis | Best 4 of 5 injections; injection vol: 150 uL; furnace temp: 680 degC; std concentration ranges: 0.35-5 mg/L C; Shimadzu TOC-L | Shimadzu TOC-L | Pacific Northwest National Laboratory
NPOC_Software_Method | NPOC_002 | MethodID_Analysis | Best 4 of 5 injections; injection vol: 50 uL; furnace temp: 680 degC; std concentration ranges: 0.35-5 mg/L C; Shimadzu TOC-L | Shimadzu TOC-L | Pacific Northwest National Laboratory
NPOC_Software_Method | NPOC_003 | MethodID_Analysis | Best 3 of 4 injections; injection vol: 100 uL; furnace temp: 680 degC; std concentration ranges: 0.5-10 mg/L C; Shimadzu TOC-L | Shimadzu TOC-L | Pacific Northwest National Laboratory
NPOC_Software_Method | NPOC_004 | MethodID_Analysis | NPOC not measured due to low sample volume. NPOC values for SPE and FTICR was based on concentration of replicates | NA | Pacific Northwest National Laboratory
NPOC_Software_Method | NPOC_005 | MethodID_Analysis | NPOC not measured in any of the replicates due to low sample volume. NPOC values for SPE and FTICR was based on the average NPOC for all the samples | NA | Pacific Northwest National Laboratory
NPOC_Software_Method | NPOC_006 | MethodID_Analysis | NPOC not measured. Vial was not filled with sample. | NA | Pacific Northwest National Laboratory
NPOC_Software_Method | NPOC_007 | MethodID_Analysis | Best 4 of 5 injections; injection vol: 150 uL; furnace temp 680 degC; std concentration ranges 0.5-50 mg/L C; Shimadzu TOC-L | Shimadzu TOC-L | Pacific Northwest National Laboratory
NPOC_Software_Method | NPOC_008 | MethodID_Analysis | NPOC measured by third party. | NA | Pacific Northwest National Laboratory
NPOC_Software_Method | NPOC_009 | MethodID_Analysis | Best 4 of 5 injections; injection vol: 150 uL; furnace temp 680 degC; std concentration ranges 0.5-20 mg/L C; Shimadzu TOC-L | Shimadzu TOC-L | Pacific Northwest National Laboratory
NPOC_Software_Method | NPOC_010 | MethodID_Analysis | Best 4 of 5 injections; injection vol: 150 uL; furnace temp 680 degC; std concentration ranges 0.35-100 mg/L C; Shimadzu TOC-L | Shimadzu TOC-L | Pacific Northwest National Laboratory
NPOC_Software_Method | NPOC_011 | MethodID_Analysis | Best 5 of 5 injections; injection col: 150 uL; furnace temp 680 degC; std concentration ranges .35-50 mg/L C; Shimadzu TOC-H |  | Pacific Northwest National Laboratory
NPOC_Software_Method | NPOC_012 | MethodID_Analysis | Best 5 of 5 injections; injection col: 150 uL; furnace temp 680 degC; std concentration ranges .35-100 mg/L C; Shimadzu TOC-H |  | Pacific Northwest National Laboratory
NPOC_Software_Method | NPOC_013 | MethodID_Analysis | Best 4 of 5 injections; injection vol: 150 uL; furnace temp 680 degC; std concentration ranges 0.35-10 mg/L C; Shimadzu TOC-H |  | Pacific Northwest National Laboratory
pH_Method | PH_000 | MethodID_Analysis | pH not measured | NA | Pacific Northwest National Laboratory
pH_Method | PH_001 | MethodID_Analysis | Measured with pH strips (range: 4-10) | NA | Pacific Northwest National Laboratory
pH_Method | PH_002 | MethodID_Analysis | Measured with Fisher Scientific Accumet Excel XL15 pH meter | Fisher Scientific Accumet Excel XL15 pH meter | Pacific Northwest National Laboratory
Randomized_Method | RAND_000 | MethodID_Analysis | Sample run order was not randomized. Replicates analyzed together in ascending order. | NA | Pacific Northwest National Laboratory
Randomized_Method | RAND_001 | MethodID_Analysis | Sample run order was randomized. | NA | Pacific Northwest National Laboratory
Randomized_Method | RAND_002 | MethodID_Analysis | Sample run order was randomized. But IC and ICP samples were run in order of ascending SpC value for all Rep 1 samples. Followed by all Rep 2 samples. Etc. | NA | Pacific Northwest National Laboratory
Randomized_Method | RAND_003 | MethodID_Analysis | Sample run order was randomized by collaborator | NA | Pacific Northwest National Laboratory
Randomized_Method | RAND_004 | MethodID_Analysis | Sample run was not randomized | NA | Pacific Northwest National Laboratory
Randomized_Method | RAND_005 | MethodID_Analysis | Sample run was not randomized. Samples analyzed together based on amount of gluteraldehyde added. | NA | Pacific Northwest National Laboratory
Randomized_Method | RAND_006 | MethodID_Analysis | Samples were run in order of ascending SpC value | NA | Pacific Northwest National Laboratory
Sediment_INC_Method | SED_INC_000 | MethodID_Analysis | Sediment was not incubated. | NA | Pacific Northwest National Laboratory
Sediment_INC_Method | SED_INC_001 | MethodID_Analysis | 10 mL of sediment and ~30 mL of unfiltered water from the site were used to incubate the sediment. Water was added to 40 mL glass vials with sediment until there was no headspace. Before incubation sediments were kept at 21degC overnight in a temperature controlled chamber prior to incubation and were aerated for ~10 min before adding the unfiltered water. The vials were placed on a shaker at 250 rpm in the dark. Dissolved oxygen concentrations were measured non-invasively with an optical oxygen meter (Fibox 3; PreSens GmbH, Germany) every 15 minutes for 2 hours or until the dissolved oxygen concentration reached 5 mg/L | NA | Pacific Northwest National Laboratory
SpC_Method | SPC_000 | MethodID_Analysis | SpC not measured | NA | Pacific Northwest National Laboratory
SpC_Method | SPC_001 | MethodID_Analysis | Samples randomized.SpC measured with VWR SpC probe. SpC was measured only for 1 rep | VWR Benchtop/Portable Conductivity Meter catalog: 23226-505 | Pacific Northwest National Laboratory
SpC_Method | SPC_002 | MethodID_Analysis | Samples randomized.SpC measured with VWR SpC probe. SpC was measured for all the samples | VWR Benchtop/Portable Conductivity Meter catalog: 23226-505 | Pacific Northwest National Laboratory
TN_Software_Method | TNSFTW_000 | MethodID_Analysis | TN not measured | NA | Pacific Northwest National Laboratory
TN_Software_Method | TNSFTW_001 | MethodID_Analysis | Best 4 of 5 injections; injection vol: 150 uL; furnace temp: 720 degC; std concentration ranges: 0.1-10 mg/L NO3-N; Shimadzu TOC-L | Shimadzu TOC-L | Pacific Northwest National Laboratory
TN_Software_Method | TNSFTW_002 | MethodID_Analysis | Best 3 of 4 injections; injection vol: 150 uL; furnace temp: 720 degC; std concentration ranges: 0.1-10 mg/L NO3-N; Shimadzu TOC-L | Shimadzu TOC-L | Pacific Northwest National Laboratory
TN_Software_Method | TNSFTW_003 | MethodID_Analysis | Best 4 of 5 injections; injection vol: 150 uL; furnace temp: 720 degC; std concentration ranges: 0.2-5 mg/L NO3-N; Shimadzu TOC-L | Shimadzu TOC-L | Pacific Northwest National Laboratory
TN_Software_Method | TNSFTW_004 | MethodID_Analysis | Best 5 of 5 injections; injection vol: 150 uL; furnace temp: 720 degC; std concentration ranges: 0.1-3 mg/L NO3-N; Shimadzu TOC-H |  | Pacific Northwest National Laboratory
XRF_Software_Method | XRF_Method_001 | MethodID_Analysis | Sample was run on a Bruker M4 plus. Detector type: XFlash 430 Window type: Custom type Detector thickness: 0.45 Si dead layer:0.029 Calibration, lin.: 10 Calibration, abs.: -955.9 Mn FWHM: 137.05 Fano factor: 0.121 |  | Environmental Molecular Sciences Laboratory/Pacific Northwest National Laboratory
Date_Received | DATEREC_000 | MethodID_DataProcessing | Date received was not recorded | NA | Pacific Northwest National Laboratory
Geochem_Method_QA_QC | Geochem_QA_000 | MethodID_DataProcessing | Geochemistry data were not evaluated for specific QA/QC requirements | NA | Pacific Northwest National Laboratory
Geochem_Method_QA_QC | Geochem_QA_001 | MethodID_DataProcessing | ICP/IC/TN/NPOC data evaluated for whether values were within +/- 10% of standard curve. If not then values replaced with "Below_Range or "Above_Range". ICP and IC data use low range and high range standard curves for data processing. Standard curves must have R-squared values >0.985. If standard curve R-squared is <0.985 then sample values are replaced with "Low_Rsq_in_STD_Curve" "Low_Rsq_in_Low_Range_STD_Curve" or "Low_Rsq_in_High_Range_STD_Curve" | NA | Pacific Northwest National Laboratory
Geochem_Method_QA_QC | Geochem_QA_002 | MethodID_DataProcessing | ICP/IC/TN/NPOC data evaluated for whether values were within +/- 10% of standard curve. If not then values replaced with "Below_Range or "Above_Range". ICP and IC data use low range and high range standard curves for data processing. Standard curves must have R-squared values >0.985. If standard curve R-squared is <0.985 then sample values are replaced with "Low_Rsq_in_STD_Curve" "Low_Rsq_in_Low_Range_STD_Curve" or "Low_Rsq_in_High_Range_STD_Curve". For TN and NPOC data: Injections with Area = 0 were dropped. If any sample had fewer than 3 injections then sample values are replaced with "Too_Few_Injections". If there were more than 3 injections - the 3 that were closest to each other were retained; the others were dropped. For each sample- the ratio of maximum to minimum Areas within the remaining 3 injections was found and if the ratio was greater than 2 - the reported concentration is given as "Injections_Too_Variable" | NA | Pacific Northwest National Laboratory
Geochem_Method_QA_QC | Geochem_QA_003 | MethodID_DataProcessing | IC/TN data evaluated for whether values were within +/- 10% of standard curve. If not then values replaced with "Below_Range or "Above_Range" and the value of the limiting standard. IC data use low range and high range standard curves for data processing. Standard curves must have R-squared values >0.985. If standard curve R-squared is <0.985 then sample values are replaced with "Low_Rsq_in_STD_Curve" "Low_Rsq_in_Low_Range_STD_Curve" or "Low_Rsq_in_High_Range_STD_Curve". For TN data: Injections with Area = 0 were dropped. If any sample had fewer than 3 injections then sample values are replaced with "Too_Few_Injections". If there were more than 3 injections - the 3 that were closest to each other were retained; the others were dropped. For each sample- the ratio of maximum to minimum Areas within the remaining 3 injections was found and if the ratio was greater than 2 - the reported concentration is given as "Injections_Too_Variable" | NA | Pacific Northwest National Laboratory
Precipitate_Description | PRECIP_000 | MethodID_Inspection | Precipitate not present | NA | Pacific Northwest National Laboratory
Precipitate_Description | PRECIP_001 | MethodID_Inspection | Precipitate present not described | NA | Pacific Northwest National Laboratory
Precipitate_Description | PRECIP_002 | MethodID_Inspection | Yellow precipitate | NA | Pacific Northwest National Laboratory
Precipitate_Description | PRECIP_003 | MethodID_Inspection | Cloudy precipitate | NA | Pacific Northwest National Laboratory
Precipitate_Description | PRECIP_004 | MethodID_Inspection | Yellow hue but no precipitate | NA | Pacific Northwest National Laboratory
Precipitate_Description | PRECIP_005 | MethodID_Inspection | Yellow hue and floating fuzzy looking substance | NA | Pacific Northwest National Laboratory
Sediment_INC_Description | SED_INC_DESC_000 | MethodID_Inspection | Domestic sediment. Incubated. | NA | Pacific Northwest National Laboratory
Sediment_INC_Description | SED_INC_DESC_001 | MethodID_Inspection | International sediment. Not incubated | NA | Pacific Northwest National Laboratory
Sediment_INC_Description | SED_INC_DESC_002 | MethodID_Inspection | Not incubated due to sample loss, shipping delays, or alternate reasons. | NA | Pacific Northwest National Laboratory
Vial_Break_Method | VB_000 | MethodID_Inspection | Vial did not break. | NA | Pacific Northwest National Laboratory
Vial_Break_Method | VB_001 | MethodID_Inspection | Vial broke at the bottom. Sample transferred to a new vial. | NA | Pacific Northwest National Laboratory
Vial_Break_Method | VB_002 | MethodID_Inspection | Vial broke on the sides. Sample was not lost. Sample transferred to a new vial. | NA | Pacific Northwest National Laboratory
Vial_Break_Method | VB_003 | MethodID_Inspection | Vial shattered. Some sample lost. Sample transferred to a new vial. | NA | Pacific Northwest National Laboratory
Vial_Break_Method | VB_004 | MethodID_Inspection | Vial seemed weak. Sample was transferred to a new vial. | NA | Pacific Northwest National Laboratory
Vial_Break_Method | VB_005 | MethodID_Inspection | Vial seemed weak. Sample was transferred to a new vial. Didn't record vial transfer at the time of transfer. | NA | Pacific Northwest National Laboratory
Vial_Break_Method | VB_006 | MethodID_Inspection | Vial did not break because vials never frozen. | NA | Pacific Northwest National Laboratory
Vial_Break_Method | VB_007 | MethodID_Inspection | Vial broke at the top. Some sample was lost. Remaining sample transferred to new vial. | NA | Pacific Northwest National Laboratory
Vial_Break_Method | VB_008 | MethodID_Inspection | Vial broke. No specifics on how/location of vial break. | NA | Pacific Northwest National Laboratory
Vial_Break_Method | VB_009 | MethodID_Inspection | No description whether vial intact or broken. | NA | Pacific Northwest National Laboratory
Vial_Break_Method | VB_010 | MethodID_Inspection | No vial sent. | NA | Pacific Northwest National Laboratory
Vial_Break_Method | VB_011 | MethodID_Inspection | Vial broke. No sample lost. Not transferred to new vial because the sample is frozen. | NA | Pacific Northwest National Laboratory
Vial_Break_Method | VB_012 | MethodID_Inspection | Vial broke but could not transfer vial because it was frozen | NA | Pacific Northwest National Laboratory
Vial_Break_Method | VB_013 | MethodID_Inspection | Vial broke. Sample lost. | NA | Pacific Northwest National Laboratory
Vial_Inspection_Method | VI_01 | MethodID_Inspection | Vial thawed during shipping | NA | Pacific Northwest National Laboratory
Vial_Inspection_Method | VI_02 | MethodID_Inspection | Vial arrived partially thawed from shipping | NA | Pacific Northwest National Laboratory
Acidification_After_Dilution_Method | ACID2_000 | MethodID_Preparation | No additional acidification after dilution (or no dilution occurred) | NA | Pacific Northwest National Laboratory
Acidification_After_Dilution_Method | ACID2_001 | MethodID_Preparation | Acidified in the lab to pH 2 after dilution (85% phosphoric acid) | NA | Pacific Northwest National Laboratory
Ammonia_Dilution Method | AMM_DILU_000 | MethodID_Preparation | Ammonia samples were not diluted | NA | Pacific Northwest National Laboratory
Ammonia_Dilution Method | AMM_DILU_001 | MethodID_Preparation | Ammonia samples were diluted using milli-q water | NA | Pacific Northwest National Laboratory
Anion_Dilution_Method | ANIONDILU_000 | MethodID_Preparation | No dilution needed | NA | Pacific Northwest National Laboratory
Anion_Dilution_Method | ANIONDILU_001 | MethodID_Preparation | Dilution into DI water to standardize SpC levels to 400 uS/cm | NA | Pacific Northwest National Laboratory
Anion_Dilution_Method | ANIONDILU_002 | MethodID_Preparation | Dilution into DI water to standardize SpC levels to 500 uS/cm | NA | Pacific Northwest National Laboratory
Anion_Dilution_Method | ANIONDILU_003 | MethodID_Preparation | Dilution not performed because SpC not measured. | NA | Pacific Northwest National Laboratory
Anion_Dilution_Method | ANIONDILU_004 | MethodID_Preparation | Dilution into DI water to standardize SpC levels to 100uS/cm | NA | Pacific Northwest National Laboratory
Anion_Dilution_Method | ANIONDILU_005 | MethodID_Preparation | Dilution into DI water to bring up to needed volume | NA | Pacific Northwest National Laboratory
Anion_Dilution_Method | ANIONDILU_006 | MethodID_Preparation | Dilution into DI water to dilute SpC level under 400 uS/cm | NA | Pacific Northwest National Laboratory
Bulk_C/N_Drying_Method | BULK_C/N_Dry_000 | MethodID_Preparation | Sample was dried in an oven at 60 degC for 144 hours |  | Pacific Northwest National Laboratory
Bulk_C/N_Drying_Method | BULK_C/N_Dry_001 | MethodID_Preparation | Sample was dried in an oven at 60 degC for 144 hours |  | Pacific Northwest National Laboratory
Bulk_C/N_Drying_Method | BULK_C/N_Dry_002 | MethodID_Preparation | Sample was dried in an oven at 60 degC for 168 hours |  | Pacific Northwest National Laboratory
Filter_Method | FLTR_000 | MethodID_Preparation | No additional filtering after field. | NA | Pacific Northwest National Laboratory
Filter_Method | FLTR_001 | MethodID_Preparation | Additional filtering needed due to particulate matter. Used Sterivex filter (0.2 um). | NA | Pacific Northwest National Laboratory
Filter_Method | FLTR_002 | MethodID_Preparation | No additional filtering because samples were frozen as sediment and extracted before use. | NA | Pacific Northwest National Laboratory
Filter_Method | FLTR_003 | MethodID_Preparation | Filtered in the lab using a 0.2 um Sterivex filter. | NA | Pacific Northwest National Laboratory
GCMS_Method | GCMS_001 | MethodID_Preparation | Samples were dried in a speedvac for 1.5 hours | NA | Pacific Northwest National Laboratory
ICR_Dilution_Method | ICRCDILU_000 | MethodID_Preparation | No dilution needed to achieve desired C concentration for FTICR | NA | Pacific Northwest National Laboratory
ICR_Dilution_Method | ICRCDILU_001 | MethodID_Preparation | Diluted in lab using DI water to achieve desired C concentration for FTICR. Dilution based on provisional NPOC data. Final NPOC data can be found in data package. | NA | Pacific Northwest National Laboratory
InorgC_Dilution_Method | INORGCDILU_000 | MethodID_Preparation | No dilution needed for DIC analysis | NA | Pacific Northwest National Laboratory
InorgC_Dilution_Method | INORGCDILU_001 | MethodID_Preparation | Dilution into DI water for DIC analysis | NA | Pacific Northwest National Laboratory
NMR_Sed_Extract_Method | NMR_SEDEXT_000 | MethodID_Preparation | NMR sediment extraction. Vortexed at level 8 for 2 hours with 1 mL of miliQ water. Centrifuged for 5 minutes at 25C and 6000g and filtered with a 0.22 um sterivex filter. | NA | Pacific Northwest National Laboratory
NMR_Sed_Extract_Method | NMR_SEDEXT_001 | MethodID_Preparation | NMR sediment extraction. Vortexed at level 8 for 2 hours with 2 mL of miliQ water. Centrifuged for 5 minutes at 25C and 6000g and filtered with a 0.22 um sterivex filter | NA | Pacific Northwest National Laboratory
NPOC_Blanks | NPOC_BLANK_000 | MethodID_Preparation | For extraction process-blanks, 20 mL of Milli-Q water were added to a proteomic friendly falcon tube; put on an Inova 2100 shaker at 375 rpms for 2 hours; and centrifuged at 6000 rcf for 5 minutes. 7 mL of the MilliQ in the tube was filtered (0.2 um Sterivex filter) and pipetted into analysis vial. | Inova 2100 shaker | Pacific Northwest National Laboratory
NPOC_Dilution_Method | NPOCDESC_000 | MethodID_Preparation | No dilution was made for NPOC analysis | NA | Pacific Northwest National Laboratory
NPOC_Dilution_Method | NPOCDESC_001 | MethodID_Preparation | Dilution was made based on protocol for NPOC analysis | NA | Pacific Northwest National Laboratory
NPOC_Dilution_Method | NPOCDESC_002 | MethodID_Preparation | Dilution was made because sample too low volume for NPOC analysis | NA | Pacific Northwest National Laboratory
NPOC_Dilution_Method | NPOCDESC_003 | MethodID_Preparation | Dilution was made because sample was too concentrated for NPOC analysis | NA | Pacific Northwest National Laboratory
Processed | PR_001 | MethodID_Preparation | Sample was not prepared. |  | Pacific Northwest National Laboratory
Sed_Extract_Method | SEDEXT_000 | MethodID_Preparation | Samples were water not sediment. No extraction needed. | NA | Pacific Northwest National Laboratory
Sed_Extract_Method | SEDEXT_001 | MethodID_Preparation | 5mL of sediment was extracted by adding 10mL of milliQ water and setting to shake on a platform shaker at 375 rpm for 2 hours. Centrifuged for 5 minutes at 25degC at 6000g then filtered with a 0.22 um sterivex filter. The filtrate was used for subsequent analyses. | NA | Pacific Northwest National Laboratory
Sed_Extract_Method | SEDEXT_002 | MethodID_Preparation | Sediment samples were extracted by adding varying amounts of sediment and milliQ water ratios then set to shake on a platform shaker for 375 rpm for 2 hours. Centrifuged for 5 minutes at 25C at 6000g then filtered with a 0.22 um sterivex filter. The filtrate was used for subsequent analyses. | NA | Pacific Northwest National Laboratory
Sed_Extract_Method | SEDEXT_003 | MethodID_Preparation | Sediment samples were extracted by adding 10 grams of sediment and 30mL of milliQ water then set to shake on a platform shaker for 375 rpm for 2 hours. Centrifuged for 5 minutes at 25C at 6000g then filtered with a 0.22 um sterivex filter. |  | Pacific Northwest National Laboratory
Sed_Extract_Method | SEDEXT_004 | MethodID_Preparation | Sediment samples were extracted by adding 5 grams of sediment and 10mL of milliQ water then set to shake on a platform shaker for 375 rpm for 2 hours. Centrifuged for 5 minutes at 25C at 6000g then filtered with a 0.22 um sterivex filter. |  | Pacific Northwest National Laboratory
SPE_Dilution_Method | NPOCDILU_000 | MethodID_Preparation | No dilution needed for FTICR-MS | NA | Pacific Northwest National Laboratory
SPE_Dilution_Method | NPOCDILU_001 | MethodID_Preparation | Dilutions into DI water for FTICR-MS | NA | Pacific Northwest National Laboratory
SPE_Dilution_Method | NPOCDILU_002 | MethodID_Preparation | Detailed dilution data can be found in the dilution csv file | NA | Pacific Northwest National Laboratory
SPE_Method | SPE_000 | MethodID_Preparation | Solid phase extraction (SPE) not performed | NA | Pacific Northwest National Laboratory
SPE_Method | SPE_001 | MethodID_Preparation | Solid phase extraction (SPE) following methods by Dittmar et al. (2008) ( https://doi.org/10.4319/lom.2008.6.230). NPOC concentration in the samples was normalized to 1.5 mg C/L prior SPE. | NA | Pacific Northwest National Laboratory
SPE_Method | SPE_002 | MethodID_Preparation | Solid phase extraction (SPE) following methods by Dittmar et al. (2008) ( https://doi.org/10.4319/lom.2008.6.230). NPOC concentration in the samples was normalized to 1.5 mg C/L prior SPE. Some of the sample volume might have been lost during elution. | NA | Pacific Northwest National Laboratory
SPE_Method | SPE_003 | MethodID_Preparation | Solid phase extraction (SPE) following methods by Dittmar et al. (2008) ( https://doi.org/10.4319/lom.2008.6.230). NPOC concentration in the samples was normalized to 1.5 mg C/L prior SPE. Stored in the -20degC for 30 days before running. | NA | Pacific Northwest National Laboratory
Subsampling_Method | SBS_001 | MethodID_Preparation | Vials thawed in cold room (4 degC). 5mL from each replicate combined into composite sample for NPOC and geochemistry. Remaining sample used for FTICR. | NA | Pacific Northwest National Laboratory
Subsampling_Method | SBS_002 | MethodID_Preparation | Vials thawed in cold room (4 degC). Subsampled each pre-acidified amber vial its own tube for NPOC analysis and used the rest for FTICR. | NA | Pacific Northwest National Laboratory
Subsampling_Method | SBS_003 | MethodID_Preparation | Vials were subsampled from unacidified vial. | NA | Pacific Northwest National Laboratory
Subsampling_Method | SBS_004 | MethodID_Preparation | Vial was not subsampled because there was not enough sample. |  | Pacific Northwest National Laboratory
Subsampling_Method | SBS_005 | MethodID_Preparation | Vial was not subsampled because the sample had a shipping delay and came up in temperature. |  | Pacific Northwest National Laboratory
Subsampling_Method | SBS_006 | MethodID_Preparation | Vial was combined with another sample to create a composite sample. |  | Pacific Northwest National Laboratory
Thawed_Method | THW_000 | MethodID_Preparation | Vials were never frozen. Stored in the dark at 4degC | NA | Pacific Northwest National Laboratory
Thawed_Method | THW_001 | MethodID_Preparation | Vials were thawed at 4 degC for 3 days | NA | Pacific Northwest National Laboratory
Thawed_Method | THW_002 | MethodID_Preparation | Vials were thawed at 4 degC overnight. | NA | Pacific Northwest National Laboratory
Thawed_Method | THW_003 | MethodID_Preparation | Vials were never frozen. Stored in the dark at 4 degC. Removed for sub-sampling or analysis. Returned to 4 degC after analysis. | NA | Pacific Northwest National Laboratory
Thawed_Method | THW_004 | MethodID_Preparation | Vials were thawed at room temp then underwater the day of processing. | NA | Pacific Northwest National Laboratory
Thawed_Method | THW_005 | MethodID_Preparation | Vials were already previously thawed, no other thawing done.
TN_Blanks | TN_BLANK_000 | MethodID_Preparation | For extraction process-blanks, 20 mL of Milli-Q water were added to a proteomic friendly falcon tube; put on an Inova 2100 shaker at 375 rpms for 2 hours; and centrifuged at 6000 rcf for 5 minutes. 7 mL of the MilliQ in the tube was filtered (0.2 um Sterivex filter) and pipetted into analysis vial. | Inova 2100 shaker | Pacific Northwest National Laboratory
TN_Dilution_Method | TNDILU_000 | MethodID_Preparation | No dilution needed to achieve total volume of 10 mL | NA | Pacific Northwest National Laboratory
TN_Dilution_Method | TNDILU_001 | MethodID_Preparation | 2X dilution (5 mL sample + 5 mL DI water) to achieve total volume of 10 mL | NA | Pacific Northwest National Laboratory
TN_Dilution_Method | TNDILU_002 | MethodID_Preparation | No dilution needed to achieve total volume of 5.5 mL | NA | Pacific Northwest National Laboratory
TN_Dilution_Method | TNDILU_003 | MethodID_Preparation | 4X dilution (1.38 mL sample + 4.12 mL DI water) to achieve total volume of 5.5 mL | NA | Pacific Northwest National Laboratory
TN_Dilution_Method | TNDILU_004 | MethodID_Preparation | 10X dilution (0.55mL sample + 4.95mL DI water) to achieve total volume of 5.5mL | NA | Pacific Northwest National Laboratory
TN_Dilution_Method | TNDILU_005 | MethodID_Preparation | 30X dilution (0.183mL sample + 5.317mL DI water) to achieve total volume of 5.5mL) | NA | Pacific Northwest National Laboratory
TN_Dilution_Method | TNDILU_006 | MethodID_Preparation | 3.6X dilution (1.5mL of sample + 4mL DI water) to achieve total volume of 5.5 mL | NA | Pacific Northwest National Laboratory
TN_Dilution_Method | TNDILU_007 | MethodID_Preparation | 6.3X dilution (.75mL of sample + 4 mL DI water) to achieve total volume of 4.75 mL | NA | Pacific Northwest National Laboratory
TN_Dilution_Method | TNDILU_008 | MethodID_Preparation | Diluted using DI to bring up to total volume of 5.5 mL | NA | Pacific Northwest National Laboratory
TN_Dilution_Method | TNDILU_009 | MethodID_Preparation | No dilution needed to achieve total volume of 5 mL | NA | Pacific Northwest National Laboratory
XRF_Method | XRF_Method_000 | MethodID_Preparation | Sample was ground on the ball mill with 1 stainless steel ball for 25 minutes at 20 beats/sec |  | Pacific Northwest National Laboratory
XRF_Method | XRF_Method_001 | MethodID_Preparation | Sample was ground on the ball mill with 1 stainless steel ball for 50 minutes at 20 beats/sec |  | Pacific Northwest National Laboratory
XRF_Method | XRF_Method_002 | MethodID_Preparation | Sample was ground on the ball mill with 2 stainless steel ball for 25 minutes at 20 beats/sec |  | Pacific Northwest National Laboratory
XRF_Method | XRF_Method_003 | MethodID_Preparation | Sample was ground on the ball mill with 2 stainless steel ball for 50 minutes at 20 beats/sec |  | Pacific Northwest National Laboratory
XRF_Method | XRF_Method_004 | MethodID_Preparation | Sample was split into 2 vials and 2 balls for 25 minutes at 20 beats/sec |  | Pacific Northwest National Laboratory
XRF_Method | XRF_Method_005 | MethodID_Preparation | Sample was split into 2 vials and 2 balls for 50 minutes at 20 beats/sec |  | Pacific Northwest National Laboratory
Acidification_Method | ACID_001 | MethodID_Preservation | Acidified in the lab to pH 2 when samples were thawed for sub-sampling (85% phosphoric acid) | NA | Pacific Northwest National Laboratory
Acidification_Method | ACID_002 | MethodID_Preservation | Vials sent to sampler were pre-acidified (10 uL 85% phosphoric acid added to 40mL empty amber borosillicate vial prior to sampling with the expectation it would be filled to 20mL). | NA | Pacific Northwest National Laboratory
Acidification_Method | ACID_003 | MethodID_Preservation | Vials sent to sampler were pre-acidified (30 uL 20% nitric acid) added to 15mL empty falcon tube prior to sampling with the expectation it would be filled to 14mL. | NA | Pacific Northwest National Laboratory
Acidification_Method | ACID_004 | MethodID_Preservation | Vials used were for in house experiment. No pre-acidification. | NA | Pacific Northwest National Laboratory
Acidification_Method | ACID_005 | MethodID_Preservation | Acidified in the lab to pH 2 using 85% phosphoric acid. | NA | Pacific Northwest National Laboratory
ICP_Dilution_Method | ICPDILU_000 | MethodID_Preservation | No dilution needed to achieve desired SpC | NA | Pacific Northwest National Laboratory
ICP_Dilution_Method | ICPDILU_001 | MethodID_Preservation | Dilution into 1.2% nitric acid. Dilution calculated to achieve 150 uS/cm based on Ca graph | NA | Pacific Northwest National Laboratory
ICP_Dilution_Method | ICPDILU_002 | MethodID_Preservation | Dilutions did occur. Exact dilution volumes not available | NA | Pacific Northwest National Laboratory
Preservative_Method | PRES_000 | MethodID_Preservation | 30 mL commercially-made RNALater added | NA | Pacific Northwest National Laboratory
Preservative_Method | PRES_001 | MethodID_Preservation | 30 mL house-made RNALater added | NA | Pacific Northwest National Laboratory
Anion_Storage_Method | ANIONST_000 | MethodID_Storage | Anions not measured | NA | Pacific Northwest National Laboratory
Anion_Storage_Method | ANIONST_001 | MethodID_Storage | Anions were run more than 1 month after collection. | NA | Pacific Northwest National Laboratory
Anion_Storage_Method | ANIONST_003 | MethodID_Storage | Anions were run within 1 month after collection. | NA | Pacific Northwest National Laboratory
Frozen_Method | FRZ_000 | MethodID_Storage | Vials were not frozen | NA | Pacific Northwest National Laboratory
Frozen_Method | FRZ_001 | MethodID_Storage | Vials were frozen (-20 degC) by the people collecting the samples. Then shipped and placed in freezer after arrriving at Pacific Northwest National Laboratory. | NA | Pacific Northwest National Laboratory
Frozen_Method | FRZ_002 | MethodID_Storage | Vials were frozen (-20 degC) after they arrived at Pacific Northwest National Laboratory. | NA | Pacific Northwest National Laboratory
Frozen_Method | FRZ_003 | MethodID_Storage | Vials were frozen (-20 degC) after they arrived at Pacific Northwest National Laboratory. Except anion vials were never frozen. | NA | Pacific Northwest National Laboratory
Frozen_Method | FRZ_004 | MethodID_Storage | Vials were frozen (-20 degC) after they arrived at Pacific Northwest National Laboratory. Vials were overfilled so they were separated into 2 vials to prevent breaking. | NA | Pacific Northwest National Laboratory
Frozen_Method | FRZ_005 | MethodID_Storage | Vials were flash frozen in liquid nitrogen then stored at -80 degC. | NA | Pacific Northwest National Laboratory
Frozen_Method | FRZ_006 | MethodID_Storage | Vials were stored at -80 degC. | NA | Pacific Northwest National Laboratory
Frozen_Method | FRZ_007 | MethodID_Storage | Vials were stored at -20 degC. | NA | Pacific Northwest National Laboratory
Frozen_Method | FRZ_008 | MethodID_Storage | Samples were frozen at -80 degC then freeze dried for 72 hours. | NA | Pacific Northwest National Laboratory
Frozen_Method | FRZ_009 | MethodID_Storage | Samples were stored at -20 degC after collection. They thawed during shipment to Sequim, they were stored at -80 degC upon receiving.