## Sheet1
Method_Name | Method_ID | Method_Type | Method_Description | Method_Instrument | Method_Lab
Anion_Typical | ANION_T_IN_000 | MethodID_Inspection | No vials broke. No precipitate or suspended particles. | NA | Pacific Northwest National Laboratory
Anion_Typical | ANION_T_ST_000 | MethodID_Storage | Stored at 4C in the dark for under 1 month after collection. | NA | Pacific Northwest National Laboratory
Anion_Typical | ANION_T_PRES_000 | MethodID_Preservation | No preservatives added. | NA | Pacific Northwest National Laboratory
Anion_Typical | ANION_T_PREP_000 | MethodID_Preparation | Detailed dilution data can be found in the anion dilution csv file | NA | Pacific Northwest National Laboratory
Anion_Typical | ANION_T_AN_000 | MethodID_Analysis | Samples were run in order of ascending SpC value. Ions: F. Cl. NO2. SO4. NO3. Br; Std concentration ranges (mg/L): (F: 0.2-10). (Cl: 0.16-80). (NO2: 0.04-20). (SO4: 0.16-80). (NO3: 0.12-60) (Br: .03-15); Dionex ICS-2000. Ion Pack AS18. prod#: 060549. serial#: 180602097. elution buffer: 22 mM KOH; suppressor current: 57 mA | NA | Pacific Northwest National Laboratory
Anion_Typical | ANION_T_DP_000 | MethodID_DataProcessing | Data evaluated for whether values were within +/- 10% of standard curve. If not then values replaced with "Below_Range or "Above_Range". ICP and IC data use low range and high range standard curves for data processing. Standard curves must have R-squared values >0.985. If standard curve R-squared is <0.985 then sample values are replaced with "Low_Rsq_in_STD_Curve" "Low_Rsq_in_Low_Range_STD_Curve" or "Low_Rsq_in_High_Range_STD_Curve". | NA | Pacific Northwest National Laboratory
Moisture_Typical | MOIS_T_IN_000 | MethodID_Inspection | Vial was intact and filled to top with sediment. | NA | Pacific Northwest National Laboratory
Moisture_Typical | MOIS_T_ST_000 | MethodID_Storage | Stored at 4C until analysis. | NA | Pacific Northwest National Laboratory
Moisture_Typical | MOIS_T_PRES_000 | MethodID_Preservation | No preservatives added. | NA | Pacific Northwest National Laboratory
Moisture_Typical | MOIS_T_PREP_000 | MethodID_Preparation | No preparation. | NA | Pacific Northwest National Laboratory
Moisture_Typical | MOIS_T_AN_000 | MethodID_Analysis | Vial was weighed then dried in 105 degC oven for 48 hours or 24 hours. If 48 hours: drying was complete after 48 hours. If 24 hours: several vials from the same oven batch were weighed again after drying; placed in the oven again for 3 hours; then weighed again. If there was no significant change to mass then analysis of all vials in oven was complete. If the mass changed then vials were put in oven again for an additional 24h until there was no mass change. | NA | Pacific Northwest National Laboratory
Moisture_Typical | MOIS_T_DP_000 | MethodID_DataProcessing | Percent moisture was calculated based on the difference between wet and dry mass. then divided by the weight of the dry or wet mass and multiplied by 100. | NA | Pacific Northwest National Laboratory
ICR_Typical | ICR_T_IN_000 | MethodID_Inspection | Vials were intact and filled to marked line. No precipitate or suspended particles. | NA | Pacific Northwest National Laboratory
ICR_Typical | ICR_T_ST_000 | MethodID_Storage | Stored at -20degC then thawed in 4degC for three days. | NA | Pacific Northwest National Laboratory
ICR_Typical | ICR_T_PRES_000 | MethodID_Preservation | 2uL of 85% phosphoric acid per 5 mL of sample was put in vial prior to adding sample. | NA | Pacific Northwest National Laboratory
ICR_Typical | ICR_T_PREP_000 | MethodID_Preparation | Prior to analysis on the FTICR-MS, samples were diluted as needed to achieve a standard NPOC concentration of 1.5 milligrams_ C_per_liter and acidified to pH 2 using 85% phosphoric acid prior solid phase extraction (SPE). If NPOC concentration of the sample was below 1.5 milligrams_C_per_liter then the sample was not diluted and 15 mL of the undiluted sample were acidified and SPEed. | NA | Pacific Northwest National Laboratory
ICR_Typical | ICR_T_AN_000 | MethodID_Analysis | Samples were run on a 12 Tesla (12T) Bruker SolariX FTICR-MS. Ion accumulation time (IAT) is appended to file name for each sample's xml files, where for example "p08" indicates an IAT of 0.08. | 12 Tesla (12T) Bruker SolariX FTICR-MS | Environmental Molecular Sciences Laboratory/Pacific Northwest National Laboratory
ICR_Typical | ICR_T_DP_000 | MethodID_DataProcessing | User completes data processing using provided instructions, Formularity, and R script. Columns in FTICR output are as follows: Mass (measured ionic mass; m/z). Columns C. H. O. N. C13. S. P. Na (atoms in the analyzed molecule. Identified by an algorithm in Kujawinski and Behn 2006 and modified as described in Tfaily et al. 2017.) El_Comp (function group(s) that are known. NeutralMass (zero-charge or neutral molecular mass. Calculated from observed ionic mass by balancing charge with protons or electrons). Error_ppm (parts per million error calculated as [expected molecular mass - measured molecular mass]/expected molecular mass * 1e6). Candidates (possible known molecules the analyzed molecule could be. Assigned by checking against a known database. Value of -1 when molecule was not checked against database). All remaining columns identified by Kit ID and Rep number (spectra for each sample where the numerical value is the peak intensity at the particular molecular mass). | NA | NA
ICR_Typical | ICR_T_IN_001 | MethodID_Inspection | Vials were intact and filled to appropriate levels. | NA | Pacific Northwest National Laboratory
ICR_Typical | ICR_T_ST_001 | MethodID_Storage | Filtered sample extract stored at 4degC for less than 72 hours before performing solid phase extraction protocol. | NA | Pacific Northwest National Laboratory
ICR_Typical | ICR_T_PRES_001 | MethodID_Preservation | No preservatives added. | NA | Pacific Northwest National Laboratory
ICR_Typical | ICR_T_PREP_001 | MethodID_Preparation | Sediment was extracted for NPOC analysis and then extracts were prepared for FTICR-MS analysis. See NPOC methods for information about extraction. Prior to analysis on the FTICR-MS, sample extracts were diluted as needed to achieve a standard NPOC concentration of 1.5 milligrams_ C_per_liter and acidified to pH 2 using 85% phosphoric acid prior solid phase extraction (SPE). If NPOC concentration of the sample extract was below 1.5 milligrams_C_per_liter then the sample extract was not diluted and 15 mL of the undiluted sample extract were acidified and SPEed. | NA | Pacific Northwest National Laboratory
ICR_Typical | ICR_T_AN_001 | MethodID_Analysis | Samples were run on a 12 Tesla (12T) Bruker SolariX FTICR-MS. Ion accumulation time (IAT) is appended to file name for each sample's xml files, where for example "p08" indicates an IAT of 0.08. | 12 Tesla (12T) Bruker SolariX FTICR-MS | Environmental Molecular Sciences Laboratory/Pacific Northwest National Laboratory
ICR_Typical | ICR_T_DP_001 | MethodID_DataProcessing | User completes data processing using provided instructions, Formularity, and R script. Columns in FTICR output are as follows: Mass (measured ionic mass; m/z). Columns C. H. O. N. C13. S. P. Na (atoms in the analyzed molecule. Identified by an algorithm in Kujawinski and Behn 2006 and modified as described in Tfaily et al. 2017.) El_Comp (function group(s) that are known. NeutralMass (zero-charge or neutral molecular mass. Calculated from observed ionic mass by balancing charge with protons or electrons). Error_ppm (parts per million error calculated as [expected molecular mass - measured molecular mass]/expected molecular mass * 1e6). Candidates (possible known molecules the analyzed molecule could be. Assigned by checking against a known database. Value of -1 when molecule was not checked against database). All remaining columns identified by same name (spectra for each sample where the numerical value is the peak intensity at the particular molecular mass). | NA | NA
NPOC_Typical | NPOC_T_IN_000 | MethodID_Inspection | Vials were intact and filled to marked line. No precipitate or suspended particles. |  | Pacific Northwest National Laboratory
NPOC_Typical | NPOC_T_ST_000 | MethodID_Storage | Stored at -20degC then thawed in 4degC for three days. |  | Pacific Northwest National Laboratory
NPOC_Typical | NPOC_T_PRES_000 | MethodID_Preservation | 2uL of 85% phosphoric acid per 5 mL of sample was put in vial prior to adding sample. |  | Pacific Northwest National Laboratory
NPOC_Typical | NPOC_T_PREP_000 | MethodID_Preparation | 5mL sample pipetted into analysis vial. |  | Pacific Northwest National Laboratory
NPOC_Typical | NPOC_T_AN_000 | MethodID_Analysis | Best 4 of 5 injections; injection vol: 150 uL; furnace temp 680 degC; std concentration ranges 0.3-20 mg/L C; Shimadzu TOC-L |  | Pacific Northwest National Laboratory
NPOC_Typical | NPOC_T_DP_000 | MethodID_DataProcessing | Data evaluated for whether values were within +/- 10% of standard curve. If not then values were replaced with "Below_Range" or "Above_Range" text. Standard curves must have R-squared values >0.985. If standard curve R-squared is <0.985 then sample values are replaced with "Low_Rsq_in_STD_Curve". Injections with Area = 0 were dropped. If any sample had fewer than 3 injections then sample values are replaced with "Too_Few_Injections". If there were more than 3 injections - the 3 that were closest to each other were retained; the others were dropped. For each sample - the ratio of maximum to minimum areas within the remaining 3 injections was found and if the ratio was greater than 2 - the reported concentration is given as "Injections_Too_Variable." |  | Pacific Northwest National Laboratory
NPOC_Typical | NPOC_T_IN_001 | MethodID_Inspection | Vials were intact and filled to appropriate levels. | N/A | Pacific Northwest National Laboratory
NPOC_Typical | NPOC_T_ST_001 | MethodID_Storage | Sediment was stored at -80C before and after lyophilization. Sediment extracts were stored at 4C in the dark for under 2 days until analyis. | N/A | Pacific Northwest National Laboratory
NPOC_Typical | NPOC_T_PRES_001 | MethodID_Preservation | No preservatives added. | N/A | Pacific Northwest National Laboratory
NPOC_Typical | NPOC_T_PREP_001 | MethodID_Preparation | Sediments were lyophilized for 3 days. Stored at -80 degC until extraction day. 10 grams of sediment and 20 mL of Milli-Q water were added to a proteomic friendly falcon tube; put on an Inova 2100 shaker at 375 rpms for 2 hours; and centrifuged at 6000 rcf for 5 minutes. 1.4 mL of the extracted supernatant was filtered (0.2 um Sterivex filter) and pipetted into analysis vial and diluted with Milli-Q water to a final volume of 7 mL. | Inova 2100 shaker | Pacific Northwest National Laboratory
NPOC_Typical | NPOC_T_AN_001 | MethodID_Analysis | Best 3 of 5 injections; injection vol: 150 uL; furnace temp 680 degC; std concentration ranges 0.25-100 mg/L C; Shimadzu TOC-L | Shimadzu TOC-L CPH/CPN with a TNM-L | Environmental Molecular Sciences Laboratory/Pacific Northwest National Laboratory
NPOC_Typical | NPOC_T_DP_001 | MethodID_DataProcessing | Data evaluated for whether values were within +/- 10% of standard curve. If not then values were replaced with "Below_Range" or "Above_Range" text. If any sample had fewer than 3 injections then sample values are replaced with "Too_Few_Injections". If there were more than 3 injections - the 3 that were closest to each other were retained; the others were dropped. For each sample - the ratio of maximum to minimum Areas within the remaining 3 injections was found and if the ratio was greater than 2 - the reported concentration is given as "Injections_Too_Variable." | N/A | Pacific Northwest National Laboratory
TN_Typical | TN_T_IN_000 | MethodID_Inspection | Vials were intact. No precipitate or suspended particles. | N/A | Pacific Northwest National Laboratory
TN_Typical | TN_T_ST_000 | MethodID_Storage | Stored at 4C in the dark until analysis. Analyzed within 1 month of collection. | N/A | Pacific Northwest National Laboratory
TN_Typical | TN_T_PRES_000 | MethodID_Preservation | No preservatives added. | N/A | Pacific Northwest National Laboratory
TN_Typical | TN_T_PREP_000 | MethodID_Preparation | 5.5mL sample pipetted into analysis vial. | N/A | Pacific Northwest National Laboratory
TN_Typical | TN_T_AN_000 | MethodID_Analysis | Best 4 of 5 injections; injection vol: 150 uL; furnace temp: 720 degC; std concentration ranges: 0.1-10 mg/L NO3-N; Shimadzu TOC-L | Shimadzu TOC-L | Pacific Northwest National Laboratory
TN_Typical | TN_T_DP_000 | MethodID_DataProcessing | Data evaluated for whether values were within +/- 10% of standard curve. If not then values were replaced with "Below_Range" or "Above_Range" text. Standard curves must have R-squared values >0.985. If standard curve R-squared is <0.985 then sample values are replaced with "Low_Rsq_in_STD_Curve". Injections with Area = 0 were dropped. If any sample had fewer than 3 injections then sample values are replaced with "Too_Few_Injections". If there were more than 3 injections - the 3 that were closest to each other were retained; the others were dropped. For each sample - the ratio of maximum to minimum Areas within the remaining 3 injections was found and if the ratio was greater than 2 - the reported concentration is given as "Injections_Too_Variable." | N/A | Pacific Northwest National Laboratory
TN_Typical | TN_T_IN_001 | MethodID_Inspection | Vials were intact and filled to appropriate levels. | N/A | Pacific Northwest National Laboratory
TN_Typical | TN_T_ST_001 | MethodID_Storage | Sediment was stored at -80C before and after lyophilization. Sediment extracts were stored at 4C in the dark for under 2 days until analyis. | N/A | Pacific Northwest National Laboratory
TN_Typical | TN_T_PRES_001 | MethodID_Preservation | No preservatives added. | N/A | Pacific Northwest National Laboratory
TN_Typical | TN_T_PREP_001 | MethodID_Preparation | Sediments were lyophilized for 3 days. Stored at -80 degC until extraction day. 10 grams of sediment and 20 mL of Milli-Q water were added to a proteomic friendly falcon tube; put on an Inova 2100 shaker at 375 rpms for 2 hours; and centrifuged at 6000 rcf for 5 minutes. 1.4 mL of the extracted supernatant was filtered (0.2 um Sterivex filter) and pipetted into analysis vial and diluted with Milli-Q water to a final volume of 7 mL. | Inova 2100 shaker | Pacific Northwest National Laboratory
TN_Typical | TN_T_AN_001 | MethodID_Analysis | Best 3 of 5 injections; injection vol: 150 uL; furnace temp: 720 degC; std concentration ranges: 0.1-3 mg/L NO3-N; Shimadzu TOC-L | Shimadzu TOC-L CPH/CPN with a TNM-L | Environmental Molecular Sciences Laboratory/Pacific Northwest National Laboratory
TN_Typical | TN_T_DP_001 | MethodID_DataProcessing | Data evaluated for whether values were within +/- 10% of standard curve. If not then values replaced with "Below_Range or "Above_Range". If any sample had fewer than 3 injections then sample values are replaced with "Too_Few_Injections". If there were more than 3 injections - the 3 that were closest to each other were retained; the others were dropped. For each sample- the ratio of maximum to minimum Areas within the remaining 3 injections was found and if the ratio was greater than 2 - the reported concentration is given as "Injections_Too_Variable" | N/A | Pacific Northwest National Laboratory
GC_Typical | GC_T_IN_000 | MethodID_Inspection |  |  | Pacific Northwest National Laboratory
GC_Typical | GC_T_ST_000 | MethodID_Storage |  |  | Pacific Northwest National Laboratory
GC_Typical | GC_T_PRES_000 | MethodID_Preservation |  |  | Pacific Northwest National Laboratory
GC_Typical | GC_T_PREP_000 | MethodID_Preparation |  |  | Pacific Northwest National Laboratory
GC_Typical | GC_T_AN_000 | MethodID_Analysis |  |  | Environmental Molecular Sciences Laboratory/Pacific Northwest National Laboratory
GC_Typical | GC_T_DP_000 | MethodID_DataProcessing |  |  | Pacific Northwest National Laboratory
LC_Typical | LC_T_IN_000 | MethodID_Inspection |  |  | Pacific Northwest National Laboratory
LC_Typical | LC_T_ST_000 | MethodID_Storage |  |  | Pacific Northwest National Laboratory
LC_Typical | LC_T_PRES_000 | MethodID_Preservation |  |  | Pacific Northwest National Laboratory
LC_Typical | LC_T_PREP_000 | MethodID_Preparation |  |  | Pacific Northwest National Laboratory
LC_Typical | LC_T_AN_000 | MethodID_Analysis |  |  | Environmental Molecular Sciences Laboratory/Pacific Northwest National Laboratory
LC_Typical | LC_T_DP_000 | MethodID_DataProcessing |  |  | Pacific Northwest National Laboratory
SpC_Typical | SpC_T_IN_000 | MethodID_Inspection | Vials were intact. No precipitate or suspended particles. |  | Pacific Northwest National Laboratory
SpC_Typical | SpC_T_ST_000 | MethodID_Storage | Stored at 4C in the dark for under 1 month after collection. |  | Pacific Northwest National Laboratory
SpC_Typical | SpC_T_PRES_000 | MethodID_Preservation | No preservatives added. |  | Pacific Northwest National Laboratory
SpC_Typical | SpC_T_PREP_000 | MethodID_Preparation | 4mL sample pipetted into analysis vial. |  | Pacific Northwest National Laboratory
SpC_Typical | SpC_T_AN_000 | MethodID_Analysis | Samples randomized. SpC measured with VWR SpC probe. SpC was measured only for 1 rep. Probe was calibrated on day of use. |  | Pacific Northwest National Laboratory
SpC_Typical | SpC_T_DP_000 | MethodID_DataProcessing | No data processing. |  | Pacific Northwest National Laboratory
pH_Typical | pH_T_IN_000 | MethodID_Inspection | [add info for field pH RC2] |  | Pacific Northwest National Laboratory
pH_Typical | pH_T_ST_000 | MethodID_Storage | [add info for field pH RC2] |  | Pacific Northwest National Laboratory
pH_Typical | pH_T_PRES_000 | MethodID_Preservation | [add info for field pH RC2] |  | Pacific Northwest National Laboratory
pH_Typical | pH_T_PREP_000 | MethodID_Preparation | [add info for field pH RC2] |  | Pacific Northwest National Laboratory
pH_Typical | pH_T_AN_000 | MethodID_Analysis | pH measured in the field using an Ultra Meter |  | Pacific Northwest National Laboratory
pH_Typical | pH_T_DP_000 | MethodID_DataProcessing | [add info for field pH RC2] |  | Pacific Northwest National Laboratory
pH_Typical | pH_T_IN_001 | MethodID_Inspection | [add info for lab pH RC2] |  | Pacific Northwest National Laboratory
pH_Typical | pH_T_ST_001 | MethodID_Storage | [add info for lab pH RC2] |  | Pacific Northwest National Laboratory
pH_Typical | pH_T_PRES_001 | MethodID_Preservation | [add info for lab pH RC2] |  | Pacific Northwest National Laboratory
pH_Typical | pH_T_PREP_001 | MethodID_Preparation | [add info for lab pH RC2] |  | Pacific Northwest National Laboratory
pH_Typical | pH_T_AN_001 | MethodID_Analysis | pH measured in the lab using XXX |  | Pacific Northwest National Laboratory
pH_Typical | pH_T_DP_001 | MethodID_DataProcessing | [add info for lab pH RC2] |  | Pacific Northwest National Laboratory
TSS_Typical | TSS_T_IN_000 | MethodID_Inspection | No bottles were broken or leaking |  | Pacific Northwest National Laboratory
TSS_Typical | TSS_T_ST_000 | MethodID_Storage | Samples were stored at 4degC |  | Pacific Northwest National Laboratory
TSS_Typical | TSS_T_PRES_000 | MethodID_Preservation | n/a |  | Pacific Northwest National Laboratory
TSS_Typical | TSS_T_PREP_000 | MethodID_Preparation | Samples were processed by filtering through a 47mm GF/F filter |  | Pacific Northwest National Laboratory
TSS_Typical | TSS_T_AN_000 | MethodID_Analysis | Need further information about analysis |  | Pacific Northwest National Laboratory
TSS_Typical | TSS_T_DP_000 | MethodID_DataProcessing | Need further information about data processing |  | Pacific Northwest National Laboratory
OCN_Typical | OCN_T_IN_000 | MethodID_Inspection | Vials were intact. No precipitate or suspended particles. |  | Pacific Northwest National Laboratory
OCN_Typical | OCN_T_ST_000 | MethodID_Storage | Stored at 4C in the dark for under 1 month after collection. |  | Pacific Northwest National Laboratory
OCN_Typical | OCN_T_PRES_000 | MethodID_Preservation | No preservatives added. |  | Pacific Northwest National Laboratory
OCN_Typical | OCN_T_PREP_000 | MethodID_Preparation | Will get info from Sammi |  | Pacific Northwest National Laboratory
OCN_Typical | OCN_T_AN_000 | MethodID_Analysis | Will get info from Sammi |  | Pacific Northwest National Laboratory
OCN_Typical | OCN_T_DP_000 | MethodID_DataProcessing | Will get info from Sammi |  | Pacific Northwest National Laboratory
RNA_Typical | RNA_T_IN_000 | MethodID_Inspection | Filters did not leak or crack.
RNA_Typical | RNA_T_ST_000 | MethodID_Storage | Stored at -80degC
RNA_Typical | RNA_T_PRES_000 | MethodID_Preservation | 2 mL of RNAlater were used to preserve the filter.
RNA_Typical | RNA_T_PREP_000 | MethodID_Preparation | Will need to get info from Reb
RNA_Typical | RNA_T_AN_000 | MethodID_Analysis | Will need to get info from Kayla?
RNA_Typical | RNA_T_DP_000 | MethodID_DataProcessing | Will need to get info from Kayla?